# PROJECT FINAL REPORT

GROUP NUMBER: T3

PROJECT NUMBER: 1-P3

TITLE: Accuracy of Brookfield and Capillary Viscometers for the Determination of Newtonian Viscosity

DATE SUBMITTED: October 10, 2002

# ROLE ASSIGNMENTS

ROLE						GROUP MEMBER

FACILITATOR				Chalothorn Vashirakovit

TIME & TASK KEEPER			Daniel-Joseph Leung

SCRIBE					Grant Foy

PRESENTER					Jeffrey Katrencik

OBJECTIVES & SPECIFIC AIMS

To quantitatively compare the Brookfield and Capillary Viscometers through the properties of sucrose solutions by:

- Determining the accuracy, precisions, and limits of both methods in finding the viscosity

- Running experiments under different temperature and concentration conditions

BACKGROUND

The response of a liquid material to shear forces can be quantified through an analysis of viscosity of the material.  Fluid materials are unable to resist shear stresses so the sucrose solution will continue to deform as long as the stress is applied.  True solutions behave in a Newtonian fashion, with viscosity independent of either stress rate and solely on the variations of temperature.  An understanding of viscosity of pure solutions can be applied to either healthy dilute blood or plasma in biological applications.

The two machines provided for analysis in the lab calculate viscosity through the use of two different shear forces. The Brookfield Viscometer applies Hook’s law to the attached spring to calculate the torque produced by the liquid between the concentric cylinders.

As either the concentration or temperature of the solution change, the torque produced on the machine changes and a different viscosity is measured.  For the Capillary Viscometers, the gravitational pull on the sucrose solution combined with the density of the solution lead to the calculation of the viscosity.  As the density of the sucrose solution changes, through a change in either the temperature or the concentration, the viscosity also changes.  Since neither of these calculations is exponential, the change in viscosity will be directly correlated and are Newtonian in effect.

THEORY AND METHODS OF CALCULATION

All Newtonian fluids demonstrate constant viscosity independent of applied shear rates, and can be calculated by the following formula (keeping a constant temperature)

τ=η(T)γ'

Sucrose solutions are expected to follow the Newtonian model.  When either temperature or concentration is varied the Newtonian model dictates that the solution will vary directly.  For the purpose of the Brookfield viscometer, the calculations are done internally.  Further research has not revealed the formula of computation.  For the capillary viscometers, the viscosity is dependant upon the density, gravitational pull (definable by time) and the capillary viscometer constant depending on size.  This calculation is defined as:

Density * CV constant * Time (sec) = Viscosity (Centipoise)

MATERIALS, APPARATUS, METHODS

- Bioengineering Laboratory III, Mitchell Litt, BE 309 Bioengineering Laboratory III Manual: "Transport Processes and Properties: Rheology,"  Fall 2002.

- Additional Materials:

  - Pipette aid

  - Plastic Volume tubes

RESULTS

Brookfield Viscometer

At a constant temperature of 20 °C, sucrose solutions of various concentrations were run under the Brookfield viscometer to determine their cP values and percentage maximum torque at varying shear rate.  From Table 4 (refer to Appendix), it shows that the Brookfield viscometer was not able to read the viscosity of any sucrose solutions with less than 0.17 g of sucrose per 1 g of water.  Below the above-mentioned concentrations, the viscometer could not reach the point whereby the cP values become constant at varying shear rate.  At the sucrose concentration of 0.17 g, 0.20 g, 0.30 g, 0.40 g, 0.50 g and 0.60 g of sucrose per 1 g of water, the average cP values are 1.62, 1.82, 2.43, 3.28, 3.82 and 4.80, respectively.  Comparing the experimentally obtained cP values to the reference viscosity of sucrose solution at 20 °C (1.57, 1.71, 2.27, 3.01, 3.99 and 5.28), the percentage errors of the values obtained were calculated.  For the sucrose solution of 0.17g, 0.20g, 0.30g, 0.40g, 0.50g and 0.60g of sucrose per 1g of water, the percentage errors of the cP values obtained are calculated to be 3.185%, 6.433%, 7.048%, 8.970%, -4.261% and –9.091%, respectively.

At a given concentration of sucrose solution, the temperatures of the solution were varied from 40-60 °C (specifically 40, 44, 48, 52, 56 and 60 °C) in order to observe how the viscosity of the sucrose solution changes with varying temperatures (refer to Table 5).  At the sucrose concentration of 0.3 g of sucrose per 1g of water, the Brookfield viscometer was not able to read the viscosity of the solution.  At the concentration of 0.4 g of sucrose per 1g of water, the Brookfield viscometer can only read the viscosity of the solution at only the temperature of 39.7 and 43.9 °C, and the cP values obtained were 1.71 and 1.59, respectively.  Compared to the reference CP values of 1.75 and 1.59, the percentage errors were calculated to be –2.286% and 0%, respectively.  The viscometer was not able to obtain the viscosity of the sucrose solution at the temperature of 48 °C and above.

At the concentration of 0.5 g of sucrose per 1 g of water, the Brookfield viscometer can only read the viscosity of the solution at only the temperature of 39.7, 43.9 and 47.6 °C, and the cP values obtained were 2.44, 1.93 and 1.71, respectively.  Compared to the reference cP values of 2.23, 2.01 and 1.85, the percentage errors were calculated to be 9.417%, -3.980% and –7.568%, respectively.  The viscometer was not able to obtain the viscosity of the sucrose solution at the temperature of 52 °C and above.

At the concentration of 0.55 g of sucrose per 1g of water, the Brookfield viscometer was able to read the viscosity of the solution at the temperature of 39.7, 43.9, 47.6, 51.6 and 59.6 °C and the cP values obtained were 2.56, 2.32, 2.11, 2.04, 1.69 and 1.54, respectively.  Compared to the reference cP values of 2.51, 2.26, 2.07, 1.89, 1.72 and 1.60, the percentage errors were calculated to be 1.992%, 2.655%, 1.932%, 7.937%, -1.744% and 3.750%, respectively.

At the concentration of 0.6 g of sucrose per 1 g of water, the Brookfield viscometer was able to read the viscosity of the solution at the temperature of 39.7, 43.9, 47.6, 51.6 and 59.6 °C, and the cP values obtained were 2.93, 2.61, 2.39, 2.08, 1.99 and 1.87, respectively.  Compared to the reference cP values of 2.82, 2.53, 2.31, 2.11, 1.92 and 1.77, the percentage errors were calculated to be 3.901%, 2.162%, 2.463%, -1.422%, 3.646% and 5.650%, respectively.

Plotting the viscosity of the sucrose solution at 20 °C against its concentration (g of sucrose per g of water), the trend of sucrose viscosity with respect to its concentration was obtained (refer to Figure 1).  The graph obtained has an R-squared value of 0.9061, and it has the slope of 6.193.

Plotting the viscosity of the sucrose solution of the concentration 0.55 g of sucrose per 1 g of water against the solution temperature (in °C), the behavior of the viscosity of the solution with respect to the solution temperature was obtained (refer to the blue line in Figure 2).  The graph has an R –squared value of 0.9833, and it has the slope of -0.0507.

Plotting the viscosity of the sucrose solution of the concentration 0.60 g of sucrose per 1 g of water against the solution temperature (in °C), the behavior of the viscosity of the solution with respect to the solution temperature was obtained (refer to the blue line in Figure 2).  The graph has an R–squared value of 0.9602, and it has the slope of -0.0535.

Capillary Viscometer

The capillary viscometer experiment was broken into two distinct sections of results. The first set of data holds temperature constant at 24.2 degrees Celsius and varies the concentration of the sucrose solution from 0 to 1 (gram sucrose / 1 gram water). The second set of data compared the same sucrose solutions 0 to 1 (gram sucrose / 1 gram water) over the variations of 40 and 50 degrees Celsius.

| 25 Degrees Celcius |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Solution (g/g) | 50 CV | 100 CV | 200 CV | Reference Value | % Error, 50 CV | % Error, 100 CV | % Error, 200 CV |
| 0 | 0.943083 | 0.923063 | 1.081792 | 0.857854815 | 9.935073831 | 7.601272204 | 26.10432096 |
| 0.11 | 1.345614 | 1.2753 | 1.42065 | 1.1897 | 13.10285885 | 7.194253377 | 19.40978747 |
| 0.25 | 1.962658 | 1.932584 | 2.04646 | 1.7482 | 12.26925812 | 10.54897946 | 17.06294077 |
| 0.43 | 4.1157 | 3.766997 | 3.701211 | 2.8528 | 44.27027919 | 32.04589011 | 29.73988545 |
| 0.67 | N/A | 4.988378 | 5.153507 | 5.2698 | N/A | -5.339667567 | -2.206164073 |
| 1 | N/A | 13.58454 | 14.05067 | 13.6514 | N/A | -0.489818494 | 2.924736634 |

Table 1 CV Viscosity at 20 °C

There are several key points to make about Table 1.  For low concentrations the 200 CV has the high error.  Since the given range for the 200 CV does not include viscosities below 2 centipoise, this error is expected.  Conversely, using the 50 CV at high concentrations would be inappropriate, both for time and accuracy.  This is evident in our .43 solution, where the percent error is almost one-half of the viscosity.  The time required to run trials of .67 or greater in the 50 CV would have required nearly an hour, returning an inaccurate viscosity reading.  Another important note is the very small percent error recorded for the .67 and 1 g/g solutions.  These two solutions were run in week four after our skills were honed, by which time most of our procedural error had been eliminated.

| 40 Degrees Celcius |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Solution (g/g) | 50 CV | 100 CV | Reference Value | % Error, 50 CV | % Error, 100 CV |
| 0.1 | 0.851665 | 0.897038 | 0.8222 | 3.580967729 | 9.0994 |
| 0.2 | 0.981023 | 1.008496 | 1.0594 | -7.39402303 | -4.8006 |
| 0.3 | 1.176791 | 1.214899 | 1.3586 | -13.38215137 | -10.5772 |
| 0.4 | 1.616217 | 1.653329 | 1.7355 | -6.871614246 | -4.7332 |
| 0.5 | 2.203562 | 2.26565 | 2.2089 | -0.240288018 | 2.5705 |
| 0.666666667 | 3.792185 | 3.373384 | 3.2774 | 15.70711893 | 2.9287 |
| 1 | 7.363406 | 6.273766 | 7.0346 | 4.674128869 | -10.8156 |

| 50 Degrees Celcius |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Solution (g/g) | 50 CV | 100 CV | Reference Value | % Error, 50 CV | % Error, 100 CV |
| 0.1 | 0.713551 | 0.755507 | 0.6902 | 3.388354364 | 9.4674 |
| 0.2 | 0.811941 | 0.851476 | 0.8782 | -7.543095493 | -3.0413 |
| 0.3 | 1.047017 | 1.100397 | 1.1105 | -5.7205917 | -0.9139 |
| 0.4 | 1.318624 | 1.410368 | 1.3973 | -5.632633234 | 0.9330 |
| 0.5 | 1.727114 | 1.774145 | 1.7504 | -1.331330276 | 1.3555 |
| 0.666666667 | 2.615769 | 2.466816 | 2.5258 | 3.56199766 | -2.3353 |
| 1 | 5.077439 | 4.768275 | 5.1101 | -0.63914503 | -6.6892 |

Table 2, 3 CV Viscosity at 40 and 50 °C, Respectively

Table 2 and 3 show that the accuracy of the CV’s are not temperature dependent from the range of 40 degrees Celsius to 50 degrees Celsius.  We were limited to testing solutions below 50 degrees because concerns about the heat resistance of the water container were raised.

DISCUSSION AND ANALYSIS

Brookfield Viscometer

In Figure 1, the equation is determined to be:

Viscosity = 7.2429 (Concentration by grams) + 0.3421

with an R-squared value of 0.994 as determined with the Brookfield Viscometer.  That is, 99.4% of data fit accordingly (within the bounds of one standard deviation) from the line.  There is a strong relationship between the viscosity and the concentration (g of sucrose / 1 gram of water).  If the concentration was “zero,” the hypothetical viscosity based on this concentration would be 0.34 cP.  Conversely, a “zero” cP would be given by a concentration of -0.047 g of sucrose / g of water.  Ideally, the cP and concentration should be “zero.”  The 95% confidence interval (0.037, 0.65) shows that the “zero” point is included in the experimental relationship (at zero concentration, 0.34 cP) determined (Table 6).  However, the standard error of 0.11 does not put the 0.34 cP within the range of acceptability, that is, 0.34 cP cannot be said to be equal to 0 cP at 0 g of sucrose / g of water.  This illustrates that the sucrose solution is not behaving ideally, with further possible error associated with the dilutions.

As a Newtonian fluid, the viscosity depends on temperature (at the given concentration).  The temperature dependence is clearly shown in Figure 2, displaying two solutions (0.552 and 0.6 g of sucrose / g of water) with R-squared values of 0.98 and 0.96, respectively.  This is an indication that there is a strong correlation between temperature and viscosity.  The equations for the relationships are:

Viscosity = -0.0507 (Temperature) + 4.5624 for 0.552 g of sucrose/ g of water

Viscosity = -0.0535 (Temperature) + 4.9726 for 0.6 g of sucrose / g of water

This shows that if the temperature was at 0 °C, then the viscosity theoretically would be 4.56 cP and 4.97 cP, respectively.  (The solution would not be frozen at 0 °C due to the freezing point depression exhibited by the introduction of a solute, the sucrose.)  Furthermore, the temperatures at which viscosity would be “zero” are 91.8 °C and 92.9 °C.  However, this cannot hold true as 0 cP is the viscosity of that of a pure, ideal water solution.

The highest rotational speed is 200 rpm and thus limits the maximum shear rate for the experiment at any concentration and temperature.  The experimental cP values were obtained by averaging the cP values at different speeds, and thus different shear rates when the % maximum torque exceeds 10%.  More consistent cP values were acquired at generally higher rotational speeds (around 120-150rpm).  If a higher rotation speed is available, more data can be obtained, and this increases the precision of the average cP values calculated.

Any cP values measured less than 1.5 were disregarded (as suggested in the Brookfield Viscometer Manual).  As seen in Table 4, solutions below 0.17 g of sucrose / g of water have discounted cP values at 20 °C.  Similarly, raising a temperature of a solution to the point where the cP is below 1.5 cP is recorded as “Not Applicable” (e.g., 0.4 g /g at 47.6 °C has an disregarded cP, Table 5).  This is a limiting factor of the Brookfield Viscometer.

Capillary Viscometer

Analysis of the viscosity curves with the CV in regards to temperature and concentration showed promising results for a Newtonian material. While the comparison to the known viscosities in Table 1, 2 and 3 all hovered about the +10 percent region, the high variance among both the CV’s at constant temperature and concentrations and amid the changing concentrations can lead to only the conclusion that poor consistency is expected when using the outdated Capillary viscometers. Additionally, the limitations on the ranges of the Capillary viscometers justify the use of a more versatile machine. As evident in the tables above, each CV is only applicable to a small subset range of viscosities.

It becomes essential to note both the times where use of the Capillary viscometer is beneficial and detrimental.  When a lab requires a relative comparative analysis between two concentrations, a small subset of trials, or trials at easily achieved temperatures, a lab technician might look towards utilizing his capillary viscometers as a viable option. Since better machinery exists, anything outside of these situations or requiring temperature either below room temperature or above 60 degrees Celsius, the investment into a better device, such as the Brookfield, becomes worthwhile.

Finally, during these trials it is suggested to alter the procedure in these three steps:

1) Use an electronic pipette aid to draw solutions through the CV.

2) Create an environment of smaller water baths to reduce time between temperature changes

3) Place solutions in some sort of water-proof container to place in water bath to

ensure sucrose solution has warmed to the displayed temperature.

CONCLUSIONS

- Comparing the accuracy of the two machines, the Brookfield Viscometer is found to be more accurate (% error ranges from -9.091 to 9.417%) than that of the Capillary Viscometer (% error ranges from -10.816 to 32.046%).

- The time required to run a trial using the Brookfield viscometer is much shorter than by using the Capillary Viscometer (approximately 90 seconds by BV compared to 600 seconds by CV).  However, the cost of the Brookfield viscometer is much higher than the cost of the Capillary viscometer (costs approximately $1000 for BV and $10 for CV).

- Comparing both the wages of individuals required to run trials on the Brookfield and the Capillary Viscometers with the essential nature of accurate viscosities, the Brookfield machine should be the choice on any lab.

REFERENCES

- Brookfield Viscometer Manual

- John Morris Scientific Pty. (2002) CANNON-Fenske Routine Viscometer Retrieved September 20, 2002 from: http://203.147.186.54/html/Cannon/JMSprod_fenskeroutineviscometer.htm

- Gilli, Roberto.  1997 Association Andrew Van Hook: “A Memento to Sugar.” Retrieved September 9. 2002 from: http://www.univ-reims.fr/Externes/AVH/MementoSugar/001.htm

APPENDIX

| g of sucrose / 1g of water | Average cP at 20C | % Max Torque at 20 °C | Reference cP at 20 °C | % error |
| --- | --- | --- | --- | --- |
| 0 | NA | NA | NA | NA |
| 0.05 | NA | NA | NA | NA |
| 0.10 | NA | NA | NA | NA |
| 0.15 | NA | NA | NA | NA |
| 0.16 | NA | NA | NA | NA |
| 0.17 | 1.62 | 6.5 | 1.57 | 3.185 |
| 0.20 | 1.82 | 8.4 | 1.71 | 6.433 |
| 0.30 | 2.43 | 17.8` | 2.27 | 7.048 |
| 0.40 | 3.28 | 16.5 | 3.01 | 8.970 |
| 0.50 | 3.82 | 19.2 | 3.99 | -4.261 |
| 0.60 | 4.80 | 22.4 | 5.28 | -9.091 |

Table 4 Brookfield data collected at the solution temperature of 200C

| g of sucrose / 1 g of water | Temperature (°C) | Average cP | % Max Torque | Reference cP | % error |
| --- | --- | --- | --- | --- | --- |
| 0.3 | 39.7 | NA | NA | NA | NA |
|  | 43.9 | NA | NA | NA | NA |
| 0.4 | 39.7 | 1.71 | 11.4 | 1.75 | -2.286 |
|  | 43.9 | 1.59 | 10.6 | 1.59 | 0.000 |
|  | 47.6 | NA | NA | NA | NA |
| 0.5 | 39.7 | 2.44 | 14.3 | 2.23 | 9.417 |
|  | 43.9 | 1.93 | 12.9 | 2.01 | -3.980 |
|  | 47.6 | 1.71 | 11.4 | 1.85 | -7.568 |
|  | 51.6 | NA | NA | NA | NA |
| 0.55 | 39.7 | 2.56 | 10.9 | 2.51 | 1.992 |
|  | 43.9 | 2.32 | 11.1 | 2.26 | 2.655 |
|  | 47.6 | 2.11 | 11.5 | 2.07 | 1.932 |
|  | 51.6 | 2.04 | 11.9 | 1.89 | 7.937 |
|  | 55.9 | 1.69 | 11.3 | 1.72 | -1.744 |
|  | 59.6 | 1.54 | 10.3 | 1.60 | -3.750 |
| 0.6 | 39.7 | 2.93 | 12.1 | 2.82 | 3.901 |
|  | 43.9 | 2.61 | 11.9 | 2.53 | 3.162 |
|  | 47.6 | 2.39 | 12.5 | 2.31 | 3.463 |
|  | 51.6 | 2.08 | 12.2 | 2.11 | -1.422 |
|  | 55.9 | 1.99 | 11.6 | 1.92 | 3.646 |
|  | 59.6 | 1.87 | 12.5 | 1.77 | 5.650 |

Table 5 Broofield data for different sucrose solution at varying temperature

Figure 1 Graph of Viscosity of sucrose solution at 20 °C

|  | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% |
| --- | --- | --- | --- | --- | --- | --- |
| Intercept | 0.342146 | 0.110065 | 3.10857 | 0.035922 | 0.036555 | 0.647737 |
| X Variable 1 | 7.242915 | 0.279732 | 25.89237 | 1.32E-05 | 6.466254 | 8.019576 |

Table 6 Statistics of Figure 1

Figure 2 Graph of Viscosity vs. Temperature of the sucrose solutions

| Date/ size | Trial | Known Density | Time | Temp | Centistrokes constant | Total sec | centistokes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1-Oct |  |  |  |  |  | reference | 0.8459 |
| 50 | 10% | 1.031391 | 3.25.97 | 40 C | 0.004 | 205.97 | 0.849742 |
| 50 | 10% | 1.031391 | 3.28.38 | 40 C | 0.004 | 208.38 | 0.859685 |
| 100 | 10% | 1.031391 | 0.58.19 | 40 C | 0.015 | 58.19 | 0.90025 |
| 100 | 10% | 1.031391 | 0.58.19 | 40 C | 0.015 | 58.19 | 0.90025 |
|  |  |  |  |  |  | reference | 1.2003 |
| 50 | 20% | 1.0755 | 3.50.41 | 40 C | 0.004 | 230.41 | 0.991224 |
| 50 | 20% | 1.0755 | 3.52.72 | 40 C | 0.004 | 232.72 | 1.001161 |
| 100 | 20% | 1.0755 | 1.02.84 | 40 C | 0.015 | 62.84 | 1.013766 |
| 100 | 20% | 1.0755 | 1.04.12 | 40 C | 0.015 | 64.12 | 1.034416 |
|  |  |  |  |  |  | reference | 1.86 |
| 50 | 30% | 1.11879 | 4.30.84 | 40 C | 0.004 | 270.13 | 1.208875 |
| 50 | 30% | 1.11879 | 4.31.13 | 40 C | 0.004 | 271.13 | 1.21335 |
| 100 | 30% | 1.11879 | 1.14.60 | 40 C | 0.015 | 74.6 | 1.251926 |
| 100 | 30% | 1.11879 | 1.14.41 | 40 C | 0.015 | 74.41 | 1.248737 |
|  |  |  |  |  |  | reference | 3.2774 |
| 50 | 40% | 1.167521 | 6.02.75 | 40 C | 0.004 | 362.75 | 1.694073 |
| 50 | 40% | 1.167521 | 6.03.89 | 40 C | 0.004 | 363.89 | 1.699396 |
| 100 | 40% | 1.167521 | 1.39.22 | 40 C | 0.015 | 99.22 | 1.737621 |
| 100 | 40% | 1.167521 | 1.39.00 | 40 C | 0.015 | 99 | 1.733768 |
|  |  |  |  |  |  | reference | 7.03 |
| 50 | 50% | 1.219935 | 8.05.35 | 40 C | 0.004 | 485.35 | 2.368382 |
| 50 | 50% | 1.219935 | 8.05.69 | 40 C | 0.004 | 485.69 | 2.370041 |
| 100 | 50% | 1.219935 | 2.12.84 | 40 C | 0.015 | 132.84 | 2.430842 |
| 100 | 50% | 1.219935 | 2.13.40 | 40 C | 0.015 | 133.4 | 2.44109 |
|  |  |  |  |  |  | reference | 0.709 |
| 50 | 10% | 1.02696 | 2.54.53 | 50 C | 0.004 | 174.53 | 0.716941 |
| 50 | 10% | 1.02696 | 2.54.12 | 50 C | 0.004 | 174.12 | 0.715257 |
| 100 | 10% | 1.02696 | 0.49.35 | 50 C | 0.015 | 49.35 | 0.760207 |
| 100 | 10% | 1.02696 | 0.49.09 | 50 C | 0.015 | 49.09 | 0.756202 |
|  |  |  |  |  |  | reference | 0.988 |
| 50 | 20% | 1.068816 | 3.12.50 | 50 C | 0.004 | 192.5 | 0.822988 |
| 50 | 20% | 1.068816 | 3.12.48 | 50 C | 0.004 | 192.48 | 0.822903 |
| 100 | 20% | 1.068816 | 0.53.81 | 50 C | 0.015 | 53.81 | 0.862695 |
| 100 | 20% | 1.068816 | 0.53.85 | 50 C | 0.015 | 53.85 | 0.863336 |
|  |  |  |  |  |  | reference | 1.491 |
| 50 | 30% | 1.11385 | 4.01.56 | 50 C | 0.004 | 241.56 | 1.076246 |
| 50 | 30% | 1.11385 | 4.02.13 | 50 C | 0.004 | 242.13 | 1.078786 |
| 100 | 30% | 1.11385 | 1.07.72 | 50 C | 0.015 | 67.72 | 1.131449 |
| 100 | 30% | 1.11385 | 1.07.84 | 50 C | 0.015 | 67.84 | 1.133454 |
|  |  |  |  |  |  | reference | 2.526 |
| 50 | 40% | 1.16233 | 4.58.12 | 50 C | 0.004 | 298.12 | 1.386055 |
| 50 | 40% | 1.16233 | 4.57.35 | 50 C | 0.004 | 297.35 | 1.382475 |
| 100 | 40% | 1.16233 | 1.22.06 | 50 C | 0.015 | 88.06 | 1.535322 |
| 100 | 40% | 1.16233 | 1.21.78 | 50 C | 0.015 | 81.78 | 1.42583 |
|  |  |  |  |  |  | reference | 5.11 |
| 50 | 50% | 1.21451 | 6.22.06 | 50 C | 0.004 | 382.06 | 1.856063 |
| 50 | 50% | 1.21451 | 6.22.41 | 50 C | 0.004 | 382.41 | 1.857763 |
| 100 | 50% | 1.21451 | 1.45.19 | 50 C | 0.015 | 105.19 | 1.916315 |
| 100 | 50% | 1.21451 | 1.44.22 | 50 C | 0.015 | 104.22 | 1.898643 |

Table 4 CV Raw Data

However, no further explanation was provided by the Brookfield Viscometer Manual as to the percent error range of having anything below 1.5 cP.